{"text_format":"3","workbook":"xlsx","n_sheets":1,"sheets":["Sheet1"],"not_rendered":{"defined_names":1},"read":"whole","date_system":"1900"}

{"sheet":"Sheet1","index":0,"n_rows":44,"n_cols":5,"rows_kept":44,"merged":[[0,0,0,4],[1,0,40,0],[41,0,46,0]]}
["血管造影高压造影剂注射系统参数及配置情况表"]
["参数","序号","详细项目（▲代表实质性需求，★代表重要参数）","产品是否符合（是/否）","备注(若不符合请在此栏填写设备相应功能)"]
[null,1,"一、主要技术规格及系统参数"]
[null,2,"1.设备用途：适用于双C臂DSA设备配套，在成人、儿童及婴幼儿患者血管造影检查中，精准控制造影剂注射，支持心血管、神经、外周血管等介入诊疗。"]
[null,3,"2.注射器头："]
[null,4,"2.1规格：单筒（或双筒可选）。"]
[null,5,"2.2显示项目：实时显示造影剂注射液量、流速、压力限值、针筒中剩余液量。"]
[null,6,"2.3针筒保温套：支持37摄氏度恒温，保障造影剂温度接近人体体温，减少对患者血管刺激。"]
[null,7,"2.4吸药速度：1-10ml/s，增量为1ml/s，满足快速吸药需求。"]
[null,8,"2.5安全保护：配备注射头位置传感器，感应吸药、排气及注射操作的正确位置，防止误操作。"]
[null,9,"2.6推杆自动回缩：卸下针筒后推杆自动回缩，提升操作安全性与便捷性。"]
[null,10,"2.7针筒装卸：针筒采用前置插入高压保护套筒，前置挡板式固定，实现便捷、快速装卸。"]
[null,11,"2.8自动吸药：具备自动吸药功能，减少人工操作误差。"]
[null,12,"3.显示控制装置："]
[null,13,"3.1显示屏：彩色LCD显示屏，支持加装双屏（分别供手术医生和技师使用），屏幕尺寸≥10英寸，显示清晰、操作便捷。"]
[null,14,"3.2控制面板：触摸屏控制，支持中文操作界面，界面简洁易懂，便于医护人员快速上手。"]
[null,15,"3.3中文操作界面：显示器控制装置支持中文显示，包括菜单、参数设置、故障提示等，满足科室医护人员操作习惯。"]
[null,16,"4.主要技术参数："]
[null,17,"4.1注射速度：0.1-45.0ml/s（单次和分阶段），增量为0.1ml/s；0.1-59.9ml/m（单次ml/m），增量为0.1ml/m，覆盖不同部位血管造影的流速需求。"]
[null,18,"4.2注射剂量：1ml-150ml，满足成人、儿童不同体重患者的剂量需求。"]
[null,19,"4.3上升/下降时间：0-9.9s，0.1s递增，可根据诊疗需求调整造影剂注射的启停速度。"]
[null,20,"4.4压力范围：100-1200psi，增量为1psi，具备压力过载保护功能，防止压力过高损伤患者血管。"]
[null,21,"4.5注射/X线延时：0.0-99.9s，增量为0.1s，可精准同步注射与DSA影像曝光，确保影像显影清晰。"]
[null,22,"4.6储存方案：可存储至少40个注射方案，涵盖不同疾病、不同部位的造影剂注射参数，便于快速调用。"]
[null,23,"4.7存储注射历史记录数：至少存储最近50次注射记录，包括注射时间、剂量、流速、压力等参数，便于术后追溯与诊疗分析。"]
[null,24,"4.8操作互锁功能：具备操作互锁功能，防止多人员同时操作导致的参数混乱。"]
[null,25,"4.9造影成像系统接口：可与双C臂DSA设备连接，实现注射和X射线曝光的联机同步，提升诊疗效率与影像质量。"]
[null,26,"4.10设备具备功能升级：支持设备功能升级，后续有新功能（如智能压力调节、远程维护）时，可通过软件或硬件升级满足临床需求。"]
[null,27,"5.注射针筒：配套150ml一次性空针筒，针筒材质符合医用标准，无菌、无热原，与注射器头兼容。"]
[null,28,"6.安装方式："]
[null,29,"▲6.1支持一体落地式双屏；"]
[null,30,"6.2可选分体床旁式、分体落地式、分体悬吊式，便于根据机房空间灵活安装。"]
[null,31,"二、配置清单/要求"]
[null,32,"1.高压注射器主机（含注射器头、显示控制装置）1台；"]
[null,33,"2.150ml一次性空针筒；"]
[null,34,"3.针筒保温套1个；"]
[null,35,"4.压力连接管2条；"]
[null,36,"三、其它："]
[null,37,"维保年限≥3年"]
[null,38,"配件维保过后只收成本费"]
[null,39,"▲服务年限过后维保费用不超过售价的10%，具体以实际谈判为准"]
["补充参数及配置\n\n（根据产品本身可自行补充，若无，可不填写）",40,null,"产品参数","备注"]
[null,41]
[null,42]