{"text_format":"3","workbook":"xlsx","n_sheets":1,"sheets":["Sheet1"],"not_rendered":{"defined_names":1},"read":"whole","date_system":"1900"}

{"sheet":"Sheet1","index":0,"n_rows":57,"n_cols":5,"rows_kept":57,"merged":[[0,0,0,4],[1,0,52,0],[53,0,57,0]]}
["新生儿辐射保暖台参数及配置情况表"]
["参数","序号","详细项目（▲代表实质性需求，★代表重要参数）","产品是否符合（是/否）","备注(若不符合请在此栏填写设备相应功能)"]
[null,1,"一、主要技术规格及系统参数"]
[null,2,"总体需求：具备辐射灯头、控制面板、婴儿床、升降装置、负压吸引装置、T-型复苏装置、空氧混合装置、黄疸治疗装置、称重装置等。"]
[null,3," 加热器材质：陶瓷材质。"]
[null,4,"温度控制模式：预热、手动、肤温三种模式。"]
[null,5,"肤温控制范围：32.0℃～38.0℃"]
[null,6," 肤温控制精度：≤±0.5℃"]
[null,7,"肤温测量范围不窄于：25.0℃～45.0℃"]
[null,8,"肤温测量精度：±0.3℃"]
[null,9," 辐射灯头支持水平/垂直方向旋转，方便临床护理。"]
[null,10,"内置负压吸引装置，吸痰压力调节范围：-140mmHg～0mmHg。"]
[null,11,"内置空氧混合装置，氧浓度调节范围：21%～100%，精度：±3%。"]
[null,12,"内置T型婴儿复苏装置，具备PIP过压保护功能。"]
[null,13,"流量调节范围：0～15L/min。"]
[null,14,"内置生命体征监护功能，提供血氧、脉率、心电等生命体征实时监测，及时发现病情变化。"]
[null,15,"主机内置黄疸治疗装置（非外挂式），高效快速退黄。"]
[null,16,"床面上有效表面内的总辐照度最大值：≥1500μW/cm2。"]
[null,17," 内置称重系统，便捷测量患儿体重，可监测体重趋势变化。"]
[null,18,"称重范围上限≥8000g，精度±10g。"]
[null,19,"≥10.4英寸彩色触摸屏，方便医护人员观察和操作。"]
[null,20," 360°报警灯显示，方便临床及时发现报警信息并处理。"]
[null,21,"挥手报警静音，非接触式的报警静音功能，保持清洁、安静的护理环境。"]
[null,22,"报警功能：具有三级声光报警，报警音量可调。"]
[null,23,"具有APGAR评分计时功能。"]
[null,24,"主机内置穿刺照明灯，亮度可调，便于检查或夜间观察。"]
[null,25,"整机高度可调，床面高度升降范围≥200mm，满足不同身高医护人员护理操作需求。"]
[null,26,"支持婴儿床倾斜角度电动调节功能，方便临床护理操作。"]
[null,27,"配备开门阻尼功能，防护门打开后无需手扶，静音落下。"]
[null,28," 内置X光拍片盒，可以在不移动婴儿的情况下拍照。"]
[null,29,"数据存储：≥24小时趋势数据存储及查看功能，方便临床跟踪/回顾。"]
[null,30," 双燕尾槽立柱设计，便于安装临床医疗器械。"]
[null,31,"配备双向抽拉式储物抽屉，方便存储。"]
[null,32,"二、配置"]
[null,33,"主机   1台"]
[null,34,"220V电源线   1根"]
[null,35,"远红外辐射加热   1个"]
[null,36,"监护模块   1个"]
[null,37,"T-组合复苏系统（包含附件）   1套"]
[null,38,"蓝光治疗仪组件   1套"]
[null,39,"称重组件   1套"]
[null,40,"热敏打印机 1台"]
[null,41,"电动升降   1个"]
[null,42,"电动调节倾斜床体   1个"]
[null,43,"阻尼门   1个"]
[null,44,"手势感应静音   1个"]
[null,45,"X光拍片盒   1个"]
[null,46,"输液架   1个"]
[null,47,"托盘   1个"]
[null,48,"三、其它："]
[null,49,"▲维保年限≥3年"]
[null,50,"▲配件维保过后只收成本费"]
[null,51,"▲服务年限过后维保费用不超过售价的10%，具体以实际谈判为准"]
["补充参数及配置\n\n（根据产品本身可自行补充，若无，可不填写）",52,null,"产品参数","备注"]
[null,53]
[null,54]
[null,null," "]