{"text_format":"3","workbook":"xlsx","n_sheets":1,"sheets":["Sheet1"],"not_rendered":{"defined_names":1},"read":"whole","date_system":"1900"}

{"sheet":"Sheet1","index":0,"n_rows":41,"n_cols":5,"rows_kept":41,"merged":[[0,0,0,4],[1,0,34,0],[35,0,40,0]]}
["一氧化氮治疗仪参数及配置参数及配置情况表"]
["参数","序号","详细项目","产品是否符合（是/否）","备注(若不符合请在此栏填写设备相应功能)"]
[null,1,"一、主要技术规格及系统参数"]
[null,2,"1.适用范围：与呼吸设备联合使用，辅助治疗肺动脉高压症"]
[null,3,"2.治疗用NO气体设备自生成，按需获取生成、传输、吸入一体连贯，无中间存储，保证NO气体活性及纯度"]
[null,4,"3.气体过滤器使用时长：≥1000小时，有效期：≥24个月"]
[null,5,"4.一氧化氮气体浓度设置范围：0ppm~80ppm，步长：1ppm 。精度：当1ppm≤NO设置值≤4ppm±0.8ppm；当 5ppm≤NO设置值≤80ppm，±2ppm,或一氧化氮设置值的10%，取大者"]
[null,6,"5.一氧化氮气体浓度监测范围：0ppm~100ppm，分辨率0.1ppm。精度：1~20ppm 范围内，± (0.5ppm+实际读数的 20%)；＞20ppm 范围内，±(0.5ppm+实际读数的 10%)"]
[null,7,"6.二氧化氮气体浓度监测范围：0ppm~10ppm，分辨率：0.1ppm，精度：±0.5ppm 或二氧化氮实际读数的±4%，取大者"]
[null,8,"7.氧气体浓度监测范围：15%~100%，分辨率：0.1%"]
[null,9,"8.采用超声氧浓度传感器监测FiO2，无需定期更换"]
[null,10,"9.主机接口：RJ45接口和USB接口，用于连接医院信息化系统和导出数据"]
[null,11,"10.具备模式包括：在不使用同步流量传感器下，治疗仪可持续送气到呼吸回路供患者治疗，且保证监测NO/NO2/O2准确性"]
[null,12,"11.整机重量（含：电池、过滤器、浓度传感器，不含台车）：≤7.5kg，方便院内外转运使用"]
[null,13,"12.监测功能包括：实时监测NO、NO2、02浓度"]
[null,14,"13.报警功能：灯光报警、声音报警、报警信息、参数闪烁"]
[null,15,"14.设备根据一氧化氮输出控制精度，对不同NO目标浓度自动设置报警上下限，并在屏幕显示上下限值；"]
[null,16,"15.波形显示包括：实时显示NO浓度趋势曲线"]
[null,17,"16.内置积水杯自动排水系统定时排水，无需人工拆除积水杯倾倒冷凝水，减少医护工作量，避免因频繁倒水操作导致漏气，适合长时间不间断治疗"]
[null,18,"17.治疗模式包括：无创/有创/麻醉/高频，主屏显示与呼吸设备连接指引图，方便医护人员快速操作使用实现治疗"]
[null,19,"18.显示要求：与NO/NO2/O2监测数据波形，同屏显示NO/NO2浓度传感器使用寿命、过滤器使用寿命，避免因更换不及时导致治疗的安全性和连续性"]
[null,20,"19.内置电池工作时间：≥120分钟，保证转运中续航能力（提供厂家说明书）"]
[null,21,"20.主屏尺寸：≥10英寸彩色TFT LCD显示"]
[null,22,"21.操作要求：全触摸屏+旋钮操作"]
[null,23,"二、配置"]
[null,24,"一氧化氮治疗仪主机 1台"]
[null,25,"台车 1台"]
[null,26,"流量传感器 1个"]
[null,27,"流量传感器电缆 1根"]
[null,28,"过滤器 1个"]
[null,29,"NO/NO2传感器组件 1套"]
[null,30,"三、其它："]
[null,31,"▲维保年限≥3年"]
[null,32,"▲配件维保过后只收成本费"]
[null,33,"▲服务年限过后维保费用不超过售价的10%，具体以实际谈判为准"]
["补充参数及配置\n（根据产品本身可自行补充，若无，可不填写）",34,null,"产品参数","备注"]
[null,35]
[null,36]
[null,37]
[null,38]
[null,39]