{"text_format":"3","workbook":"xlsx","n_sheets":1,"sheets":["Sheet1"],"not_rendered":{"defined_names":1},"read":"whole","date_system":"1900"}

{"sheet":"Sheet1","index":0,"n_rows":43,"n_cols":5,"rows_kept":43,"merged":[[0,0,0,4],[1,0,36,0],[37,0,42,0]]}
["射频消融治疗仪参数及配置情况表"]
["参数","序号","详细项目","产品是否符合（是/否）","备注(若不符合请在此栏填写设备相应功能)"]
[null,1,"一、主要技术规格及系统参数"]
[null,2,"1.射频输出频率：455-490 kHz"]
[null,3,"2.目标温度：50-120℃，增量≤1℃"]
[null,4,"3. 温度范围：控制50-120℃，增量≤1度，测量0-124℃，增量≤1℃，精度 ≤±3℃"]
[null,5,"4.控制模式：治疗过程中可以通过功率控制和温度控制两种系统进行操作。在“功控”模式下运作时，操作者可以设定目标时间和功率。设备会自动地在设定的时间内保持此功率。在“温控”模式下运作时，操作者可以设定最大允许的输出功率，目标时间和目标温度，设备会自动地调节功率输出（在最大允许值之内）以达到和保持设定的目标温度。"]
[null,6,"5.在整个治疗过程，操作者可随时调节功率以达到所需的治疗效果。"]
[null,7,"6.输出功率：5-150W，增量≤1W，精度 ≥±10%。"]
[null,8,"7.阻抗监测范围：10 -500Ω。"]
[null,9,"8.主机可实时显示电极周围组织的阻抗值和温度。"]
[null,10,"9.系统控制：只有在阻抗范围大于10Ω和小于500Ω时工作。"]
[null,11,"11.有自动调温、控温功能，能根据消融灶组织的温度，自动调整热能，确保整个靶区达到靶温。"]
[null,12,"12.系统需具备智能识别保护装置，在误操作和设备不正常时自动切断射频输出以保证患者的安全。"]
[null,13,"13.显示功能：LED显示与功能设定，需具有实时温度、功率、阻抗显示功能；能设定目标温度和功率、设定时间和倒记时间，以及信息显示；特定 LED 显示RF 开启和错误。"]
[null,14,"14.电极：具有测温装置可以测量并反馈组织温度，能有效控制靶区温度，有高强度带形状记忆功能的可伸缩子电极。能配置多爪电极（探针可伸展2cm、3cm、4cm、5cm，工作长度≤10cm、≤15cm，套管外经15G或17G电极直接穿刺进针）和单针电极。和多针电极"]
[null,15,"15.电源工作参数：单相交流110-230V和50/60Hz。"]
[null,16,"16.连接：电源线，脚踏开关，电极连接线，回路连接线。"]
[null,17,"17.具有预警功能，警示灯与声音报警功能。"]
[null,18,"18.安装于金属、双层、四轮台车上，底面采用螺栓固定。"]
[null,19,"19. 电气安全保护：符合中华人民共和国国标及安全要求。"]
[null,20,"21.同时具有单针、爪针和水冷电极"]
[null,21,"22.水冷单针和多针电极最长长度为25cm"]
[null,22,"23.电极针为钛合金，表面有一层纳米涂层减少碳化，利于吸收"]
[null,23,"24.温度辐射范围可控。使用更安全"]
[null,24,"二、配置"]
[null,25,"系统主机 1台"]
[null,26,"电极连接线 1根"]
[null,27,"回路连接线  2根"]
[null,28,"电源线  1根"]
[null,29,"脚踏开关 1个"]
[null,30,"负极板 2个"]
[null,31,"四轮台车 1台"]
[null,32,"三、其它："]
[null,33,"▲维保年限≥3年"]
[null,34,"▲配件维保过后只收成本费"]
[null,35,"▲服务年限过后维保费用不超过售价的10%，具体以实际谈判为准"]
["补充参数及配置\n（根据产品本身可自行补充，若无，可不填写）",36,null,"产品参数","备注"]
[null,37]
[null,38]
[null,39]
[null,40]
[null,41]