{"text_format":"3","workbook":"xlsx","n_sheets":1,"sheets":["Sheet1"],"not_rendered":{"defined_names":1},"read":"whole","date_system":"1900"}

{"sheet":"Sheet1","index":0,"n_rows":31,"n_cols":5,"rows_kept":31,"merged":[[0,0,0,4],[1,0,24,0],[25,0,30,0]]}
["医用血液冷冻冰箱参数及配置情况表"]
["参数","序号","详细项目","产品是否符合（是/否）","备注(若不符合请在此栏填写设备相应功能)"]
[null,1,"一、主要技术规格及系统参数"]
[null,2,"1、温度范围-10°C～-30°C可调节；"]
[null,3,"2、微电脑控制，LCD数码显示箱内温度，显示精度0.1℃；"]
[null,4,"3、具有多种故障报警：高温报警、低温报警、传感器故障报警、环温高报警、断电报警、门开报警；"]
[null,5,"4、具有多种报警方式：声音蜂鸣报警、数字闪烁报警、符号闪烁报警，远程报警接口；"]
[null,6,"5、多重保护功能：开机延时保护、停机间隔保护、显示面板密码保护、断电记忆数据保护、传感器故障保护运行；"]
[null,7,"6、具有断电报警功能，且在产品断电后能有数字温度显示≥24小时；"]
[null,8,"7、采用HC环保制冷剂和制冷系统，明确制冷剂用量，制冷剂用量符合国家安全标准,可燃制冷剂不能高于150g；"]
[null,9,"8、根据低温保存箱国家标准GB/T 20154要求，低温保存箱铭牌上要标注制冷剂的详细名称及装入量；"]
[null,10,"9、采用碳氢高效压缩机，更节能环保;"]
[null,11,"0、德国EBM风机；"]
[null,12,"11、设定-30℃的特性点温度均匀性±3℃，全温区温度均匀性±5℃以内；"]
[null,13,"12、立式双门结构，优化嵌入式双密封条设计，密封保温效果好，环保节能；"]
[null,14,"13、内藏式蒸发器设计，金属喷粉内胆，防腐蚀，制冷快；"]
[null,15,"14、箱内大空间，搁架间距上下可调，便于存放不同尺寸的物品；"]
[null,16,"15、门锁+锁鼻一体式手把门锁设计，即可一把钥匙一把锁，又可外加挂锁，可随意配置任意挂锁，实现多人管理；"]
[null,17,"16、具有2个测试孔，方便实验使用和监控箱内温度；"]
[null,18,"17、脚轮+底脚设计，便于移动和锁定；"]
[null,19,"20、配备温度监控模块。"]
[null,20,"二、其它："]
[null,21,"▲维保年限≥3年"]
[null,22,"▲配件维保过后只收成本费"]
[null,23,"▲服务年限过后维保费用不超过售价的10%，具体以实际谈判为准"]
["补充参数及配置\n（根据产品本身可自行补充，若无，可不填写）",24,null,"产品参数","备注"]
[null,25]
[null,26]
[null,27]
[null,28]
[null,29]