{"text_format":"3","workbook":"xlsx","n_sheets":1,"sheets":["Sheet1"],"not_rendered":{"defined_names":1},"read":"whole","date_system":"1900"}

{"sheet":"Sheet1","index":0,"n_rows":39,"n_cols":5,"rows_kept":39,"merged":[[0,0,0,4],[1,0,32,0],[33,0,38,0]]}
["真空超声清洗消毒器参数及配置情况表"]
["参数","序号","详细项目","产品是否符合（是/否）","备注(若不符合请在此栏填写设备相应功能)"]
[null,1,"一、主要技术规格及系统参数"]
[null,2,"1、适用范围：适用于医院手术室、消毒供应室对硬式内镜、软式管腔器械以及常规外科手术器械、诊疗器械、器具进行全自动预清洗、真空清洗、漂洗、消毒上油和干燥等；"]
[null,3,"2、清洗容积：≥110L"]
[null,4,"3、应具备全自动清洗、漂洗、消毒上油和干燥功能，采用可编程序控制器对整个过程进行全自动控制；"]
[null,5,"4、预清洗、清洗、漂洗、超声、上油、干燥等每个功能能够单独运行，能够一机多用，针对普通手术器械、外来器械、精密器械有专门的预清洗程序，可代替手工刷洗；"]
[null,6,"5、具备真空超声波清洗技术（可提供拥有相关技术的证明资料）；"]
[null,7,"6、具有多级变压脉冲或减压沸腾清洗技术，非其他真空清洗技术（可提供拥有相关技术的证明资料）；"]
[null,8,"7、具有真空干燥技术（可提供拥有相关技术的证明资料）；"]
[null,9,"8、具备腔镜专用灌流清洗快速接头，既保证管腔内清洗和干燥质量又符合规范要求,（可提供腔镜专用灌流清洗快速接头的实物图片）；"]
[null,10,"9、具有预清洗功能，减少二次污染机率，一个流程结束即可灭菌，无需手工再干燥；"]
[null,11,"10、具备高低液位安全控制及故障自动检测报警系统；"]
[null,12,"11、有自动降温排水系统，排水更安全；"]
[null,13,"12、数据端口免费对外开放，可实现与信息化管理追溯系统相连接，数据可保存3年以上；"]
[null,14,"13、采用彩色触摸屏人机界面，智能数显实时超声波功率、频率、温度、时间、上油时间、干燥时间等数据；"]
[null,15,"14、不需要专用耗材，耗材对外开放"]
[null,16,"15、加热方式：电加热，非蒸汽加热；"]
[null,17,"16、开关门方式：脚踏式电动翻盖门，具备红外线智能防压装置（可提供脚踏式开关和红外线智能防压装置的实物图片）；"]
[null,18,"17、清洗数量：≥40个腔镜专用灌流清洗快速接头；"]
[null,19,"18、场地受限设备最大外形尺寸:≤1300mm；"]
[null,20,"19、标准工作周期：≤60min；"]
[null,21,"20、超声波频率：40KHZ；"]
[null,22,"21、设备总功率:≤25KW；"]
[null,23,"二、配置要求："]
[null,24,"2.1、主机1台"]
[null,25,"2.2、专用清洗篮筐2个"]
[null,26,"2.3、配件篮4个"]
[null,27,"2.4、牵引管16根"]
[null,28,"三、其它："]
[null,29,"▲维保年限≥3年"]
[null,30,"▲配件维保过后只收成本费"]
[null,31,"▲服务年限过后维保费用不超过售价的10%，具体以实际谈判为准"]
["补充参数及配置\n（根据产品本身可自行补充，若无，可不填写）",32,null,"产品参数","备注"]
[null,33]
[null,34]
[null,35]
[null,36]
[null,37]