{"text_format":"3","workbook":"xlsx","n_sheets":1,"sheets":["Sheet1"],"not_rendered":{"defined_names":1},"read":"whole","date_system":"1900"}

{"sheet":"Sheet1","index":0,"n_rows":29,"n_cols":5,"rows_kept":29,"merged":[[0,0,0,4],[1,0,25,0],[26,0,31,0]]}
["射频消融发生器参数及配置情况表"]
["参数","序号","详细项目（▲代表实质性需求，★代表重要参数）","产品是否符合（是/否）","备注(若不符合请在此栏填写设备相应功能)"]
[null,1,"一、主要技术规格及系统参数"]
[null,2,"1.工作电源:AC 220V-240V 52HZ"]
[null,3,"2.最大外形尺寸:50cmx40cmx15cm"]
[null,4,"3.重量:10-15KG"]
[null,5,"与射频消融钳配合使用时:"]
[null,6,"1.双回路，交替输出射频能量；实时检测组织阻抗并显示；40秒自动停止射频能力输出；"]
[null,7,"2.工作频率:450-480kHz"]
[null,8,"3.输出功率:25-30W，110-120Ω"]
[null,9,"4.输出电压:50-60Vrms"]
[null,10,"5.透壁提示:400Ω"]
[null,11,"与射频消融笔配合使用时："]
[null,12,"1.双回路，交替输出射频能量；实时检测组织阻抗并显示；40秒自动停止射频能量输出；"]
[null,13,"2.工作频率:450-480kHz"]
[null,14,"3.输出功率:12-15W"]
[null,15,"4.输出电压:75-80Vrms"]
[null,16,"5.升级配备连接口；标测系统连接口"]
[null,17,"二、配置"]
[null,18,"射频消融发生器  1台"]
[null,19,"脚踏开关   1个"]
[null,20,"电源线   1根"]
[null,21,"三、其它："]
[null,22,"▲维保年限≥3年"]
[null,23,"▲配件维保过后只收成本费"]
[null,24,"▲服务年限过后维保费用不超过售价的10%，具体以实际谈判为准"]
["补充参数及配置\n\n（根据产品本身可自行补充，若无，可不填写）",25,null,"产品参数","备注"]
[null,26]
[null,27]